## Slide 1
SK3889 Luminescence Spectroscopy of Semiconductors: Theory and Experiment
(7.5 credits)
To pass the course:
one approved presentation (figures can be downloaded from the course page)
60% mark on home assignments (6 tasks total)
laboratory work (design and conduct an experiment using our equipment and,
for example, your samples)
Our equipment:
micro-PL (imaging and spectroscopy, including Fabry-Perot interferometry)
low-temperature micro-PL
PLE (with a tunable excitation source)
time-resolved measurements with APD
quantum yield measurements with an integrating sphere
tunable light source based on laser-driven Xe lamp
NIR/visible absorption in the integrating sphere…

## Slide 2
Main definitions. Specific case of semiconductor luminescence
Photoluminescence (PL) and photoluminescence excitation (PLE) techniques
Photodetectors: photomultiplier tube (PMT), avalanche photodiode (APD), CCD
Monochromators and spectrographs: resolution, throughput and signal-to-noise ratio
Spectral and resolution corrections
Other methods: lock-in amplifier, photon counting and Fourier spectroscopy
Outline of the first lecture

## Slide 3
Main definitions. Specific case of semiconductor luminescence
1. Thermal radiation. Black-body source.
Wien’s law: l [nm]= 2.9·106/T[K], human thermal radiation peak is ~ 10 mm
Thermally provided energy gives rise to the emission
THIS IS NOT LUMINESCENCE
2. Scattering and reflected radiation.
Raman scattering (interaction of initial photons with phonons)
Rayleigh scattering (interaction of inital photons with particles smaller than l0)
Brillouin scattering (interaction of initial photons with refractive index variations)
These processes are fast and almost elastic (very small energy loss for initial photons)

## Slide 4
Chart: 
| photo | electro | chemo | bio | cathodo | mechano | thermo | x-ray | sono | tribo
LUMINESCENCE | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
”slow” relaxation of extra energy provided to the system through light emission
(”slow” compared to light wave oscillations (fs): from nanoseconds to hours)
Main definitions. Specific case of semiconductor luminescence
Classified by the source of extra energy:
Classified by the decay rate:
Chart: 
| fluorescence (fast) | phosphorescence (slow)
LUMINESCENCE | 1 | 1
LUMINESCENCE:

## Slide 5
Main definitions. Specific case of semiconductor luminescence
For a solid-state object luminescence can be caused by the material itself or by impurities
Low-temperature luminescence spectra of Si crystal:
I – intrinsic components
P, Al ,B, Sb, As – common impurities/dopants
NP – no-phonon line
TA, LO, TO – specific phonon modes (transverse,
longitudinal, acoustic and optical)
EHL – electron-hole liquid
BE, FE – bound and free excitons
LUMINESCENCE for impurity analysis
A – nominally pure Si crystal
B – doped with Sb and As
C – doped with B and As
Peak position, intensity and linewidth all contain
information about physical processes in the sample

## Slide 6
Main definitions. Specific case of semiconductor luminescence
Peak identification can be ambiguous sometimes, especially for broad lines
Intrinsic spectra
Extrinsic spectra

## Slide 7
Main definitions. Specific case of semiconductor luminescence
Low-temperature measurements give a better picture – sharper lines, higher intensity:
linewidth depends on interaction with phonons, less phonons at low-temperature
thermal quenching can reduce luminescence intensity (e.g. dissociation of excitons)
Room temperature CL of ZnS with
1. Zn 2. Cu 3. Ag 4. Au 5. Mn
Low temperature CL of ZnO:
sharp excitonic transitions
One can use information from known samples to identify peaks correctly

## Slide 8
Main definitions. Specific case of semiconductor luminescence
Photoluminescence (PL) and photoluminescence excitation (PLE) techniques
Photodetectors: photomultiplier tube (PMT), avalanche photodiode (APD), CCD
Monochromators and spectrographs: resolution, throughput and signal-to-noise ratio
Spectral and resolution corrections
Other methods: lock-in amplifier, photon counting and Fourier spectroscopy
Outline

## Slide 9
Photoluminescence (PL) and photoluminescence excitation (PLE) techniques
Luminescence is almost independent on the way the external energy is supplied
Photoluminescence is a practical, convenient method of luminescence measurements:
configuration (c) with a dichroic mirror is the most typical one
General geometrical configuration:
fixed excitation and spectrally dispersed emission Ilum(lem) for lex0 – PL
fixed detection wavelength under variable excitation Ilum(lex) for lem0 - PLE

## Slide 10
Photoluminescence (PL) and photoluminescence excitation (PLE) techniques
PL setup examples
laser as an excitation source, lex = 488 nm
sample in the cryostat
excitation and collection through the same
optical path with a dichroic filter – bright field
Hg lamp with a filter as an excitation source
sample in the cryostat
collection through a different path than excitation -
dark field configuration
mechanical beam chopper for lock-in amplification
Radiation must be focused on the entrance slit of monochromator

## Slide 11
Photoluminescence (PL) and photoluminescence excitation (PLE) techniques
PLE setup example
- for weakly absorbing sample:
IPLE(hnex)=IPL(hnem)/I0(hnex) ~ a (hnex)
PLE gives an absorption spectrum
PLE is an alternative to transmission measurements,
for non-transparent substrates or weak absorption samples
also gives you best excitation wavelength
for PL measurements
Stokes shift between PL and PLE peaks contains
information about material structure (luminescence via
defect centers, for example)
It is related to the density of states DOS
through the oscillator strength f12:
IPLE(hnex) = DOS*f12

## Slide 12
Main definitions. Specific case of semiconductor luminescence
Photoluminescence (PL) and photoluminescence excitation (PLE) techiques
Photodetectors: photomultiplier tube (PMT), avalanche photodiode (APD), CCD
Monochromators and spectrographs: resolution, throughput and signal-to-noise ratio
Spectral and resolution corrections
Other methods: lock-in amplifier, photon counting and Fourier spectroscopy
Outline

## Slide 13
Photodetectors: photomultiplier tube (PMT), avalanche photodiode (APD), CCD
PMT is a single channel detector (vacuum tube):
photoelectron emission from cathode with
a low work function φ < hn
electron is accelerated to the first dynode
secondary electron emission with coefficient
d (typically 3-10) multiplies signal
voltage drop across array of n dynodes drives
and multiplies electrons to the anode
final number of electrons (gain) G = dn  106
PMT operation principle

## Slide 14
Photodetectors: photomultiplier tube (PMT), avalanche photodiode (APD), CCD
PMT main parameters
spectral range (defined by photocathode material, typically 150 – 1100 nm)
gain G  105-7
quantum efficiency (photon-to-electron conversion probability) ηph ~ 0.1-20 %
anode sensitivity (output current per light power unit) k(l) ~ 103-5 A/W
maximum current I ~ 10-100 mA not to overheat and damage electrodes
PMT background and noise
dark current Id (thermionic emission of electrons from photocathode) can be reduced by cooling of the PMT with a Peltier element or LN2
the main problem is not the background itself but the noise associated with it
lowest detectable signal level (noise equivalent power) NEP ~ 10-15-16 W
Shot noise
random nature of the incident photon flux and photoemission
at any given time the output signal value fluctuates around <S>
and obeys Poisson statistics
shot noise: N =S1/2

## Slide 15
Photodetectors: photomultiplier tube (PMT), avalanche photodiode (APD), CCD
APD operation principle
APD is a single channel detector (solid state):
photon absorption generates an electron-hole pair (Eg < hn)
the pair is dissolved in p-i-n junction under reverse bias
carriers are accelerated to the electrodes
impact ionization creates an avalanche of carriers
photovoltage is created at the detector output
APD main parameters
spectral range (defined by the bandgap of active layer, e.g. 600-1800 nm for Ge)
ionization coefficient (number of ionization events per unit distance) ae or ah
practically only one type of carriers is used for signal, so ionization ratio ae/ah → 0 or ∞
photodiode gain G = ae·w, w – width of the multiplication region
for fast response a small size of the sensitive area is used ~ 1 mm2 (downto 20x20 μm)

## Slide 16
Photodetectors: photomultiplier tube (PMT), avalanche photodiode (APD), CCD
Comparison of single channel detectors
PMT
APD
smaller, more robust
(solid state, no cooling)
lower dark noise
higher detectivity (1/NEP)
larger sensitive area (~3x12 mm2)
suitable for spectroscopy
fast response suitable for
single photon counting
fast response suitable for
single photon counting

## Slide 17
Single photon counting detectors: PMT, APD, hybrids, multi-channel plates,…
Bulter, A., Single-Photon Counting Detectors for the Visible Range Between 300 and 1,000 nm.
In Advanced Photon Counting, Kapusta, P., Ed. Springer: 2015.
Buller G., Single-Photon Detectors for Infrared Wavelengths in the Range 1–1.7 μm
In Advanced Photon Counting, Kapusta, P., Ed. Springer: 2015.
QD-based detectors
Superconducting detectors
Upconversion-based detectors

## Slide 18
Photodetectors: photomultiplier tube (PMT), avalanche photodiode (APD), CCD
CCD history
Nobel prize in physics 2009: ”…In 1969 Willard S. Boyle and George E. Smith invented
the first successful imaging technology using a digital sensor, a CCD (Charge-Coupled
Device). The CCD technology makes use of the photoelectric effect, as theorized by
Albert Einstein and for which he was awarded the 1921 year's Nobel Prize.
By this effect, light is transformed into electric signals. The challenge when designing
an image sensor was to gather and read out the signals in a large number of image
points, pixels (= picture elements), in a short time.”
CCD is a multi-channel solid state detector
(not just an array of many photodiodes!)

## Slide 19
CCD operation principle
Photodetectors: photomultiplier tube (PMT), avalanche photodiode (APD), CCD
generation of electron-hole pairs by incident light
electrons are stored in a potential well below varying bias
electrodes
collected charge is transferred row-by-row to the
readout register by gradual variation of individual biases
ampification of the collected signal, visualization
pixel size is 20x20 mm, 12x12 mm…
array size is 1024x256, 1024x1024,…
Read-out is not from every pixel, as in photodiode arrays
but from a common register – read-out noise reduction

## Slide 20
CCD main parameters
Photodetectors: photomultiplier tube (PMT), avalanche photodiode (APD), CCD
spectral sensitivity (defined by detector bandgap, e.g. for Si 400-1000 nm)
multiplexing advantage: parallel acquisition from n channels reduces signal-to-noise
ratio by n1/2 compared to serial read-out, like in PMTs (1/n faster total readout)
spectral coverage (range of wavelength detectable with one shot) from 30 to 300 nm
quantum efficiency is wavelength dependent
dark current is very low for cooled detectors
total noise consists of shot noise, dark current
noise (suppressed by cooling) and readout noise
(constant)
signal grows linearly with time, while shot noise
as a square root – longer acquisition time can yield
S/N>1

## Slide 21
Main definitions. Specific case of semiconductor luminescence
Photoluminescence (PL) and photoluminescence excitation (PLE) techiques
Photodetectors: photomultiplier tube (PMT), avalanche photodiode (APD), CCD
Monochromators and spectrographs: resolution, throughput and signal-to-noise ratio
Spectral and resolution corrections
Other methods: lock-in amplifier, photon counting and Fourier spectroscopy
Outline

## Slide 22
Monochromators and spectrographs: resolution, throughput and signal-to-noise ratio
monochromator focuses signal with a
specific wavelength at the output slit for a
single channel detector
spectrograph produces spectrum focused
at the output port for a multichannel detector
imaging spectrograph has better optics for
aberration free spectrum image
They all contain a light dispersion element
(prism or grating) and focal planes coincide with
input and output ports
Different geometries, Czerny-Turner type (d)
is the most common nowadays
Definitions and types

## Slide 23
Monochromators and spectrographs: resolution, throughput and signal-to-noise ratio
Dispersion - how far different wavelengths end up from each other at the output plane
linear dispersion: L = f·dθ/dl = dℓ/dl ,
dθ/dl – angular dispersion defined by grating (prism)
f – focal length the monochromator/spectrometer
typical reciprocal linear dispersion L-1 ~ 1-10 nm/mm
(for a CCD chip 26 mm long and L-1 = 5 nm/mm the total
spectral range detected at a single exposure is 130 nm)
Resolving power – ability to distinct between two closely spaced lines
R = l/l, l – resolution (resolvable difference between 2 lines)
limited by:
- diffraction of the light beam at the entrance slit;
- dispersion element strength:
- size (for a prism); groove density (for a grating);
l = L-1 · l, l is a slit width (defines source size)
- small source size and large dispersive element/high groove density – best spectral resolution

## Slide 24
Monochromators and spectrographs: resolution, throughput and signal-to-noise ratio
Optimal slit width (source size)
slit diffraction angle d = l/ l
there is an optimal value of the slit lopt
below which the incoming light misses the
collimating lens due to diffraction on the slit
practical values for l are 10-100 mm
thus optimal resolution R corresponds to this value
Optimal dispersive element
holographic gratings (made by two laser interference) are usually used these days;
several gratings mounted on a rotating turrent provide different resolution and
working spectral range;
Caution: gratings have several order diffractions – should be taken into account
Caution: spectral response of gratings is wavelength dependent

## Slide 25
Monochromators and spectrographs: resolution, throughput and signal-to-noise ratio
Throughput of monochromators
the output photon flux (for broad spectrum): φ= ,
k – attenuation, B – entrance slit brightness,  - angular dispersion,
= h/f – angular height of the slit, R (resolution) = l/l.
Throughput grows:
quadratically with slit width (open slit more, but resolution drops)
linearly with angular slit height (fill the full slit with the source)
quadratically with dispersive object size (larger prisms or gratings)
Throughput of spectrographs
the outpux flux normalized by illuminated area of the detector (irradiance): Eλ=
h/h0 – relative height of the illuminated part of the detector
Throughtput grows:
linearily with slit width;
linearily with fraction of detector area illuminated;
quadratically with aperture ratio W/f (proper spectrograph design)

## Slide 26
Monochromators and spectrographs: resolution, throughput and signal-to-noise ratio
Signal-to-noise ratio in a scanning monochromator
D*
A – detector sensitivity area,
– scanning time of for a spectrum feature,
τ= – response time of the detection system, <<1
D* - normalized detectivity = D , D = 1/NEP
grows as square root with acquisition time
grows linearly with entrance slit brightness B
Light collection system should be matched with
spectrometer optics for best performance:
=
Lens L1 should be with minimized aberrations
and matching lens L2 to avoid losses

## Slide 27
Outline
Main definitions. Specific case of semiconductor luminescence
Photoluminescence (PL) and photoluminescence excitation (PLE) techiques
Photodetectors: photomultiplier tube (PMT), avalanche photodiode (APD), CCD
Monochromators and spectrographs: resolution, throughput and signal-to-noise ratio
Spectral and resolution corrections
Other methods: lock-in amplifier, photon counting and Fourier spectroscopy

## Slide 28
Spectral and resolution corrections
Even the best spectrometer/detector system distorts the signal
non-uniform spectral response of all optical components: system correction function Q (l):
  can be measured with a black body radiation source of known temperature
  all measured broad spectra must be corrected with Q (l)
  especially important for the edges of detection sensitivity
photon wavelength-to-energy correction
  simplest approach is hn[eV]=1239/l[nm], not perfectly correct
    for linear in l monochromators and all spectrographs:
    I(hn) = dN/d(hn)= l2 dN/dl = l2 I(l)
    for linear in hn monochromators: I(l) = I(hn)/ l2
    Simple approach can be used as a first approximation in most cases
Spectral corrections

## Slide 29
Spectral and resolution corrections
Even the best spectrometer/detector system distorts the signal
Resolution corrections
A known source size (slit width) can be used to deconvolute true emission line:
  measured signal i(E0)=-)d
  s(E0-E) – apparatus function (gives linewidth broadening)
  s(E0-E) ~ exp(-(E0-E)2/E’2)
this function can be measured with a source of atomic lines:
the dominant broadening is the one introduced by the spectrometer
can be neglected for very broad spectra
can be taken into account for narrow spectra assuming Gaussians:
l2meas=l’2+ l2real

## Slide 30
Outline
Main definitions. Specific case of semiconductor luminescence
Photoluminescence (PL) and photoluminescence excitation (PLE) techiques
Photodetectors: photomultiplier tube (PMT), avalanche photodiode (APD), CCD
Monochromators and spectrographs: resolution, throughput and signal-to-noise ratio
Spectral and resolution corrections
Other methods: lock-in amplifier, photon counting and Fourier spectroscopy

## Slide 31
Other luminescence detection methods
Lock-in amplifier
Enhancement of weak signals using modulation-synchronized amplification:
- signal e1 is modulated with a mechanical chopper
- external modulation source e2 is synchronized with it
- as a result only frequency-matching signal is amplified
in a mixer device e3 = e1 ∙e2
- DC output signal exists only for phase-matched signal
f1 = f2 after long RC integration time.
Small shifts in phase can be easily detected
  this is a tuning parameter for achieving maximum
  amplification compensating for delays in lines of two signals

## Slide 32
Other luminescence detection methods
Lock-in amplifier: limitations
not suitable for luminescence with long lifetimes
(weak signal in the “window” between excitation pulses)
(very fast PL decay in (a) and very slow decay in (d))
one should modulate emission beam instead, if possible
not suitable for extremely weak signals, where
discrete nature of photons becomes apparent
(large number of photons in (a), very low in (c) – zero
output due to the lack of continuous component)

## Slide 33
Other luminescence detection methods
Photon counting
PMT, APD and CCD can work in photon counting mode
the result is the pulse train of incoming photons vs. time
Important parameters for PMTs:
pulses with fast response time (< 10 ns)
high quantum efficiency (> 20%)
narrow response counts distribution
proper operation voltage (voltage independent
plateau region)

## Slide 34
Other luminescence detection methods
Fourier luminescence spectroscopy
Applied for high spectral resolution (l<0.001 nm) spectroscopy (in the infrared region)
Michelson interferometer with a varying arm length
Temporal interference signal is Fourier processed
Multiplexing (all wavelengths at once) and
high throughput (no slit necessary) are advantages
- High cost and limited application range are disadvantages